## COA
CERTIFICATE OF ANALYSIS
| CERTIFICATE No: |  | CAN/2016/
PRODUCT: | BENZYLPENICILLIN G SODIUM INJECTION BP 1 MEGA |  |  | REF NO: NDQA201509297
DATE RECEIVED | LABELCLAIM | Each vial contains Benzylpenicillin Sodium BP 1,000,000 I.U.(600mg)
29.09.2015
BATCH No. 586140462 | PRESENTATION | White free flowing powder for intramuscular or intravenous administration, contained in a clear colourless glass vial fitted with a grey rubber closure and an aluminum band and 50 such vials packed in a unit box.
MFG. DATE | MANUFACTURER | CSPC Zhongnuo Pharmaceuticals (Shijiazhuang) Co., Ltd.
Apr. 2014
EXP. DATE: | ADDRESS | No. 47 Fengshou Road, Shijiazhung, Hebei Province, CHINA.
Apr. 2017
CLIENT REF NO | CLIENT | Pharmacy and Poisons Board P.O. Box 27663-00506 Nairobi, KENYA
-
BNZP/05/11.03.2015/00139 | TEST(S) REQUESTED | Uniformity of Weight, Sterility, Identification & Assay
RESULTS
TEST | METHOD | COMPENDIA | SPECIFICATION | DETERMINED | REMARKS
Uniformity of Weight | Weight | BP 2016 Vol. V App. XII G. | < 2 vial contents deviate by more than 10% from mean vial content weight | None Deviates | COMPLIES
Sterility | Membrane Filtration | BP 2016 Vol. V App. XVI A. | No Microbial Growth | No Microbial Growth | COMPLIES
Identification | HPLC | USP 38 NF 33 Page 4797 | RT of the Component Peak in the assay sample preparation corresponds to that in the assay standard preparation | Super-imposable peak at RT 6.3 -/+ 10% min. present in both the assay sample and standard preparations | COMPLIES
Assay | HPLC | USP 38 NF 33 Page 4797 | 90.0 - 120.0% | 111.2% (RSD = 1.7% ; n = 7) | COMPLIES
CONCLUSION : | The sample complies with the specifications for the tests performed.
ANALYST: | MR. F. NAULA |  | DATE: 17-11-2015
ANALYST: | MR. M. SANGALE |  | DATE: 03-03-25016
REVIEWER: | DR. E TANUI |  | DATE: 21-03-2016
DIRECTOR: | HEZEKIAH CHEPKWONY |  | DATE: 04-18-2016