## COA
CERTIFICATE OF ANALYSIS
| CERTIFICATE No: |  | CAN/2016/
PRODUCT: | TENOFOVIR DISOPROXIL FUMARATE/ LAMIVUDINE/ EFAVIRENZ TABLETS 300 mg/300 mg /600 mg |  |  | REF NO: NDQB201604851
DATE RECEIVED | LABELCLAIM | Each film coated tablet contains: Tenofovir disoproxil fumarate 300 mg, Lamivudine USP 300 mg and Efavirenz USP 600 mg.
07.04.2016
BATCH No. 3048664 | PRESENTATION | Off white coloured, caplet-shaped, biconvex faced tablets, embossed 'M 152' on one face and plain on the other, packed in a white plastic multidose container carrying 30 tablets in a unit box.
MFG. DATE | MANUFACTURER | MYLAN Laboratories Limited.
Dec. 2015
EXP. DATE: | ADDRESS | F-4 & F-12, MIDC, Malegaon, Sinnar, Nashik-422 113, Maharashtra, INDIA
Nov. 2017
CLIENT REF NO | CLIENT | Kenya Medical Supplies Authority, P.O. Box 47715 - 00100, Nairobi, KENYA
-
- | TEST(S) REQUESTED | Uniformity of Weight, Identification, Dissolution & Assay
RESULTS
TEST | METHOD | COMPENDIA | SPECIFICATION | DETERMINED | REMARKS
Uniformity of Weight | Weight | BP 2016 Vol. V App XII C | NMT 2 tablets deviate by more than 5% from mean tablet weight | None Deviates | COMPLIES
Identification | HPLC | Adopted In-House Method | RT of the component peaks in the assay sample preparation correspond to those in the assay standard preparation | Super-imposable peaks at RT 5.4, 15.1 and 21.9 -/+ 10% min. present in both the assay sample and standard preparations | COMPLIES
Dissolution | HPLC | Adopted In-House Method | Lamivudine and Tenofovir disoproxil fumarate No tablet less than 85% (n=6; Efavirenz No tablet less than 80% (n=6) | Lamivudine 102.0% (RSD=1.6%; n=6) Efavirenz 97.0% (RSD=1.6%; n=6) Tenofovir disoproxil fumarate 101.0% (RSD=2.0%; n=6) | COMPLIES COMPLIES COMPLIES
Assay | HPLC | Adopted In-House Method | 90.0 - 110.0% | Lamivudine 101.8% (RSD=1.2%; n=9) Efavirenz 95.8% (RSD=0.7%; n=9) Tenofovir disoproxil fumarate 101.8% (RSD=1.7%; n=9) | COMPLIES COMPLIES COMPLIES
CONCLUSION : | The sample complies with the specifications for the tests performed.
ANALYST: | DR. E. MUTUA |  | DATE: 10-05-2016
REVIEWER: | DR. N. MWAURA |  | DATE: 23-05-2016
ANALYST: | EMMANUEL TANUI |  | DATE: 05-30-2016