## COA
CERTIFICATE OF ANALYSIS
| CERTIFICATE No: |  | CAN/2016/
PRODUCT: | Diclofenac Sodium Injection |  |  | REF NO: NDQD201507032
DATE RECEIVED | LABELCLAIM | Each mL contains: Diclofenac Sodium B.P. 25 mg. Benzly Alcohol B.P. 4.0 %v/v & Water for Injection B.P. q.s.
21.07.2015
BATCH No. FEDSI-002 | PRESENTATION | Clear colourless solution in a 3 mL colourless glass ampoule, packed in a blister strip of 10 ampoules in a unit box.
MFG. DATE | MANUFACTURER | LABORATE Pharmaceuticals India Ltd.
Jan.2015
EXP. DATE: | ADDRESS | Rajban Road,Paonta Sahib, H.O.:E-11,INDL.AREA,PANIPAT-132 103, INDIA.
Dec.2017
CLIENT REF NO | CLIENT | Kenya Medical Supplies Authority, P.O. BOX 47715-00100, Nairobi, KENYA
-
- | TEST(S) REQUESTED | Sterility, Identification, Assay & Acidity/Alkalinity
RESULTS
TEST | METHOD | COMPENDIA | SPECIFICATION | DETERMINED | REMARKS
Sterility | Membrane Filtration | BP 2012 Vol. V App. XVI A. | No Microbial Growth | No Microbial Growth | COMPLIES
Identification | HPLC | Adopted In- House Method | RT of the Major Peak in the assay sample preparation corresponds to that in the assay standard preparation | Super-imposable peak at RT 7.3 Ãƒâ€šÃ‚Â± 10% min. present in both the assay sample and standard preparations. | COMPLIES
Assay | HPLC | Adopted In- House Method | 95.0 - 105.0% | 100.9% (RSD = 0.4% ; n = 9) | COMPLIES
Acidity/Alkalinity | HPLC | Adopted In- House Method | 8.2-9.0 | 8.5 | COMPLIES
CONCLUSION : | The sample complies with the specifications of the tests performed.
ANALYST: | MR. E. NGAMAU |  | DATE: 20-08-2015
ANALYST: | DR. L. WANGARI |  | DATE: 02-12-2015
REVIEWER: | DR. E. TANUI |  | DATE: 21-08-2015
REVIEWER: | DR. G. WANGANGA |  | DATE: 09-12-2015
ANALYST: | ERNEST MBAE |  | DATE: 12-21-2015
DIRECTOR: | HEZEKIAH CHEPKWONY |  | DATE: 01-12-2016