## COA
CERTIFICATE OF ANALYSIS
| CERTIFICATE No: |  | CAN/2015/
PRODUCT: | HEPTAVIR - 150 TABLETS |  |  | REF NO: NDQD201508124
DATE RECEIVED | LABELCLAIM | Each film coated tablet contains: Lamivudine USP 150 mg.
06.08.2015
BATCH No. E141003 | PRESENTATION | Brown colored, caplet shaped, biconvex faced tablets, single scored on one face and embossed 'H' and 'V' on either side of the score and '150' on the opposite unscored face, packed in a white plastic multi-dose container carrying 60 tablets in a unit box.
MFG. DATE | MANUFACTURER | HETERO Labs Limited.
Jun. 2014
EXP. DATE: | ADDRESS | 22 - 110, I.D.A., Jeedimetla, Hyderabad - 500 055, INDIA.
May 2017
CLIENT REF NO | CLIENT | National AIDS and STI Control Programme P.O. BOX 19361 Nairobi
-
- | TEST(S) REQUESTED | Uniformity of Weight, Identification, Dissolution & Assay
RESULTS
TEST | METHOD | COMPENDIA | SPECIFICATION | DETERMINED | REMARKS
Uniformity of Weight | Weight | B.P. 2012 Vol. V App XII C | NMT 2 tablets deviate by more than 5% from mean weight | None Deviates | COMPLIES
Identification | HPLC | Adopted In-House Method | RT of the major peak in the assay sample preparation corresponds to that in the assay standard preparation | Super-imposable peak at RT 2.9 Â± 10% min. present in both the assay sample and standard preparations | COMPLIES
Dissolution | HPLC | Adopted In-House Method | No tablet less than 80.0% [n=6] | 97.9% (RSD = 4.9%; n = 6) | COMPLIES
Assay | HPLC | Adopted In-House Method | 90.0 - 110.0% | 103.7% (RSD = 0.7%; n = 9) | COMPLIES
CONCLUSION : | The product complies with the specifications for the tests performed.
ANALYST: | DR. E. MUTUA |  | DATE: 01-10-2015
REVIEWER: | DR. G. WANG'ANG'A |  | DATE: 19-10-2015
DIRECTOR: | HEZEKIAH CHEPKWONY |  | DATE: 11-05-2015