## COA
CERTIFICATE OF ANALYSIS
| CERTIFICATE No: |  | CAN/2016/
PRODUCT: | ASAQ WINTHROP 100 mg/270 mg TABLETS |  |  | REF NO: NDQD201508224
DATE RECEIVED | LABELCLAIM | Each tablet contains: Artesunate 100 mg & Amodiaquine Base 270 mg (as hydrochloride 352.64 mg)
24.08.2015
BATCH No. 4MA236 | PRESENTATION | Yellow and off white coloured, bi-layered flat faced tablets, single scored and embossed "A" & "S" on either side of the score on the off white face and embossed "100" on the yellow face, packed in an aluminium foil blister strip of 6 tablets and 25 such strips in a box.
MFG. DATE | MANUFACTURER | MAPHAR
Jun.2014
EXP. DATE: | ADDRESS | Km 10,route cotiere 111, (SEE COMMENTS ON DRAFT PRINT OUT) Zenata Ain Sebaa, Casablanca, MOROCCO.
May.2017
CLIENT REF NO | CLIENT | Pharmacy Board Of Sierra Leone, Central Medical Stores, Free Town, SIERRA LEONE
-
PVG/GFM/MAK-PUB/18/3/15 | TEST(S) REQUESTED | Uniformity of Weight, Uniformity of Content, Identification, Dissolution & Assay
RESULTS
TEST | METHOD | COMPENDIA | SPECIFICATION | DETERMINED | REMARKS
Uniformity of Weight | Weight | BP 2016 Vol. IV App. XII C | Not more than 2 tablets deviate by more than 5% of the tablet weight. | None Deviates | COMPLIES
Identification | HPLC | Adopted In - House Method | RT of the Major peaks in the assay sample preparation correspond to those in the assay standard preparation. | Super-imposable peaks at RT 2.0 & 8.2 -/+ 10% min. present in both the assay sample and standard preparations | COMPLIES
Dissolution | HPLC HPLC | Adopted In - House Method | After 30 Mins No tablet is less than 80% | Amodiaquine HCl 85.0% (RSD = 1.4%; n =6) Artesunate 96.9%; (RSD = 0.7%; n=6) | COMPLIES COMPLIES
Assay | HPLC HPLC | Adopted In - House Method | 90.0 - 110.0% | Amodiaquine HCl 96.6% (RSD = 1.6%; n =9) Artesunate 99.6% (RSD = 1.2%; n=9) | COMPLIES COMPLIES
CONCLUSION : | The sample complies with the specifications of the tests performed.
ANALYST: | MICHAEL SANGALE |  | DATE: 06-01-2016
REVIEWER: | DR.WANG'ANG'A |  | DATE: 06-01-2016
ANALYST: | ERNEST MBAE |  | DATE: 01-11-2016
DIRECTOR: | HEZEKIAH CHEPKWONY |  | DATE: 01-25-2016