## COA
CERTIFICATE OF ANALYSIS
| CERTIFICATE No: |  | CAN/2015/
PRODUCT: | EFAVIRENZ 600 mg TABLETS |  |  | REF NO: NDQD201510425
DATE RECEIVED | LABELCLAIM | Each film-coated tablet contains Efavirenz 600 mg.
08.10.2015
BATCH No. 7224450 | PRESENTATION | Beige coloured, caplet-shaped, biconvex faced tablets, plain on both faces, packed in a white plastic multidose container carrying 30 tablets in a unit box.
MFG. DATE | MANUFACTURER | STRIDES Arcolab Limited.
May. 2015
EXP. DATE: | ADDRESS | S-No. 36/7, Suragajakkanahalli, Indlavadi Cross, Anekal Taluk, Bangalore - 562 106, INDIA.
Apr. 2018
CLIENT REF NO | CLIENT | Christa Cepuch c/o MSF -
-
- | TEST(S) REQUESTED | Uniformity of Weight, Identification, Dissolution & Assay
RESULTS
TEST | METHOD | COMPENDIA | SPECIFICATION | DETERMINED | REMARKS
Uniformity of Weight | Weight | B.P. 2012 Vol. V App XII C | NMT 2 tablets deviate by more than 5% from mean weight | None Deviates | COMPLIES
Identification | HPLC | Adopted In-House Method | RT of the major peak in the assay sample preparation corresponds to that in the assay standard preparation | Super-imposable peaks at RT 25.6 +/- 10% min. present in both the assay sample and standard preparations | COMPLIES
Dissolution | HPLC | Adopted In-House Method | No tablet less than 80.0% [n=6] | 98.3% (n=6; RSD=1.0%) | COMPLIES
Assay | HPLC | Adopted In-House Method | 90.0 % - 110.0% | 102.1%; (n=9; RSD=1.1%) | COMPLIES
CONCLUSION : | The sample complies with the specifications of the tests perfomed.
ANALYST: | DR. S. MWANGI |  | DATE: 02-11-2015
REVIEWER: | DR. G. WANG’ANG’A |  | DATE: 04-11-2015
ANALYST: | NICHOLAS MWAURA |  | DATE: 11-05-2015
DIRECTOR: | HEZEKIAH CHEPKWONY |  | DATE: 11-09-2015