## COA
CERTIFICATE OF ANALYSIS
| CERTIFICATE No: |  | CAN/2016/
PRODUCT: | FAST 250mg SUPPOSITORIES |  |  | REF NO: NDQD201511541
DATE RECEIVED | LABELCLAIM | Each suppository contains Paracetamol BP 250 mg.
10.11.2015
BATCH No. S0024062 | PRESENTATION | Off white coloured bullet shaped suppositories for rectal use only packed in a blister strip of 5 suppositories and 2 such strips in a unit box
MFG. DATE | MANUFACTURER | The ACME Laboratories Ltd.
Nov. 2014
EXP. DATE: | ADDRESS | Dhaka, BANGLADESH.
Nov. 2016
CLIENT REF NO | CLIENT | Highridge Pharmaceuticals Ltd. P. O. Box 32982 - 00600, Nairobi, KENYA.
-
- | TEST(S) REQUESTED | Uniformity of Weight, Microbial Load, Identification, Disintegration & Assay
RESULTS
TEST | METHOD | COMPENDIA | SPECIFICATION | DETERMINED | REMARKS
Uniformity of Weight | Weight | BP 2016 Vol. V App XII C | NMT 2 suppositories deviate by more than 5% from mean weight | None Deviates | COMPLIES
Microbial Load | Plate Count | USP 38 NF 33 Pg 2008 | Total Aerobic Microbial Count (TAMC) NMT 200 CFU/g Total Combined Yeasts /Moulds Count (TYMC) NMT 20 CFU/g | < 10 TAMC CFU/g < 10 TYMC CFU/g | COMPLIES COMPLIES
Identification | HPLC | USP 38 NF 33 Pg 2008 | RT of the component peak in the assay sample preparation corresponds to that in the assay standard preparation | Super-imposable peak at RT 4.3 -/+ 10% min. present in both the assay sample and standard preparations | COMPLIES
Disintegration | Disintegration | B.P. 2016 Vol. V App XII A | Not more than 30 minutes [n=6] | 16 min. 40 sec. | COMPLIES
Assay | HPLC | USP 38 NF 33 Pg 2008 | 90.0 - 110.0% | 92.6% (RSD=0.3%; n=9) | COMPLIES
CONCLUSION : | The sample complies with the specifications for the tests performed.
ANALYST: | MR. F. NAULA |  | DATE: 30-11-2015
ANALYST: | DR. E. MUTUA |  | DATE: 01-02-2016
REVIEWER: | DR. E. TANUI |  | DATE: 15-02-2016
ANALYST: | ERNEST MBAE |  | DATE: 02-16-2016
DIRECTOR: | HEZEKIAH CHEPKWONY |  | DATE: 02-22-2016