## COA
CERTIFICATE OF ANALYSIS
| CERTIFICATE No: |  | CAN/2015/347
PRODUCT: | Paracetamol Tablets 500 mg |  |  | REF NO: NDQ20150505000
DATE RECEIVED | LABELCLAIM | Each Tabltet contains Paracetamol 500 mg
29.05.2015
BATCH No. TTYB50 | PRESENTATION | A blister pack of 10 in a unit box
MFG. DATE | MANUFACTURER | MYLAN LABS LTD
May.2015
EXP. DATE: | ADDRESS | F-4 & F-12, MIDC, Malegaon, Sinnar, Nashik-422 113, Maharashtra, INDIA
Sept.2015
CLIENT REF NO | CLIENT | Chemonics Kenya Pharma LLC P.o Box 1325 00606 Nairobi, Kenya
-
| TEST(S) REQUESTED | Uniformity of Weight, Identification, Dissolution & Assay
RESULTS
TEST | METHOD | COMPENDIA | SPECIFICATION | DETERMINED | REMARKS
Uniformity of Weight | Weight | B.P. 2012 Vol .V App XII C | NMT 2 tablets deviate by more than 5% from mean weight | None Deviates | COMPLIES
Identification | HPLC | MOA 113 | Super-imposable peak at RT 7.0 -/+ 10% min. present in both the assay sample and standard preparations | RT of the major peak in the assay sample preparation corresponds to that in the assay standard preparation | COMPLIES
Dissolution | HPLC | MOA 113 | NLT 80% | Paracetamol 96%(RSD = 0.97% ; n = 6) | COMPLIES
Assay | HPLC | MoA 13339 | 90 - 110% | Paracetamol 99%(RSD = 0.57% ; n = 12) | COMPLIES
CONCLUSION : | The sample complies with the specifications of the tests perfomed.
ANALYST: | MS. M MAGDA |  | DATE: 05-29-2015
REVIEWER: | DR. E MBAE |  | DATE: 05-29-2015
DIRECTOR: | DR. H. CHEPKWONY |  | DATE: 05-29-2015