**CERTIFICATE OF ANALYSIS**

**CERTIFICATE No:** CAN/2016-2017/39

|  |  |  |  |
| --- | --- | --- | --- |
| **PRODUCT** | COSATRIMÂ® SUSPENSION | | **REF. NO: NDQB2016061171** |
| **DATE RECEIVED:** 2016-06-09 | **LABEL CLAIM:** | Each 5 mL contains: Sulphamethoxazole BP 200 mg and Trimethoprim BP 40 mg. | |
| **BATCH NO.:** 60655 | **PRESENTATION:** | Pink coloured suspension for oral administration in a 100 mL amber coloured plastic bottle and supplied along with a 15 mL graduated clear plastic measuring cap, packed in a unit box. | |
| **MFG. DATE:** May 2016 | **MANUFACTURER:** | COSMOS Limited. | |
| **EXP. DATE:**  Apr. 2019 | **ADDRESS:** | Rangwe Rd; Off Lunga Lunga Road, P.O. Box 41433 - 00100, Nairobi, KENYA. | |
| **CLIENT REF NO:** | **CLIENT:** | Kenya Medical Supplies Authority P.O. Box 47715 - 00100, Nairobi, KENYA. | |
|  | **TEST(S) REQUESTED:** | Microbial Load, Identification, Assay, Acidity/Alkalinity | |

**RESULTS**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **TEST** | **METHOD** | **COMPEDIA** | **SPECIFICATION** | **DETERMINED** | **REMARKS** |
| **Microbial Load** | Microbial Load | BP 2016 Vol V App XVI B | Total Aerobic Microbial Count (TAMC) NMT 200 CFU/mL | TAMC < 10 CFU/mL | **COMPLIES** |
| **Microbial Load** |  |  | Total Combined Yeast/Moulds Count (TYMC) NMT 20 CFU/mL | TYMC < 10 CFU/mL |  |
| **Identification** | HPLC | USP 38 NF 33 Page 5397 | RT of the component peaks in the assay sample preparation correspond to those in the assay standard preparation | Super-imposable peaks at RT 3.5 & 6.9 ± 10% min. present in both the assay sample and standard preparations | **COMPLIES** |
| **Assay** | HPLC | USP 38 NF 33 Page 5397 | 90.0 - 110.0% | Trimethoprim 96.9% (RSD=0.2%; n=9) | **COMPLIES** |
| **Assay** |  |  |  | Sulphamethoxazole 93.9% (RSD=0.1%; n=9) |  |
| **Acidity/Alkalinity** | pH | USP 38 NF 33 Page 5397 | 5.0 - 6.5 | 5.9 | **COMPLIES** |

**CONCLUSION:** The product complies with the specifications for the tests performed.

|  |  |  |  |
| --- | --- | --- | --- |
| **ANALYST:** | MICHAEL SANGALE | .......................................................... | DATE: 06-21-2016 |
| **ANALYST:** | FRANCIS NAULA | .......................................................... | DATE: 06-21-2016 |
| **ANALYST:** | GEORGE WANG'ANG'A | .......................................................... | DATE: 07-18-2016 |
| **ANALYST:** | ERNEST MBAE | .......................................................... | DATE: 07-25-2016 |
| **DIRECTOR:** | HEZEKIAH CHEPKWONY | .......................................................... | DATE: 08-02-2016 |