**CERTIFICATE OF ANALYSIS**

**CERTIFICATE No:** CAN/2016-2017/113

|  |  |  |  |
| --- | --- | --- | --- |
| **PRODUCT** | SULFRAN DS TABLETS | | **REF. NO: NDQB201608081** |
| **DATE RECEIVED:** 2016-08-23 | **LABEL CLAIM:** | Each tablets contains: Sulfamethoxazole B.P. 800 mg and Trimethoprim B.P 160 mg. | |
| **BATCH NO.:** 621462 | **PRESENTATION:** | Off white coloured, caplet shaped, biconvex faced tablets, single scored on one face and plain on the other, packed in a blister strip of 10 tablets and 10 such strips packed in a unit box. | |
| **MFG. DATE:** Jul. 2016 | **MANUFACTURER:** | UNIVERSAL Corporation Ltd. | |
| **EXP. DATE:**  Jun. 2019 | **ADDRESS:** | Club Road, Plot No. 13777, P.O Box 1748 - 00902, Kikuyu, KENYA. | |
| **CLIENT REF NO:** | **CLIENT:** | Kenya Medical Supplies Authority, P.O. Box 47715 - 00100, Nairobi, Kenya. | |
|  | **TEST(S) REQUESTED:** | Uniformity of Weight, Identification, Dissolution, Assay | |

**RESULTS**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **TEST** | **METHOD** | **COMPEDIA** | **SPECIFICATION** | **DETERMINED** | **REMARKS** |
| **Uniformity of Weight** | Weight | BP 2016 Vol. V App. XII C | NMT 2 tablets deviate by more than 5% from mean weight | None Deviates | **COMPLIES** |
| **Identification** | HPLC | USP 38 NF 33 page 5399 | RT of the major peaks in the assay sample preparation correspond to those in the assay standard preparation | Super-imposable peaks at RT 3.8 and 5.2 -/+ 10% min. present in both the assay sample and standard preparations | **COMPLIES** |
| **Dissolution** | HPLC | USP 38 NF 33 page 5399 | No tablet less than 75% (n = 6) | Sulfamethoxazole Average = 97%, Range | **COMPLIES** |
| **Dissolution** |  |  |  | 94 - 98% (RSD=2.1%, n=6) |  |
| **Dissolution** |  |  |  | Trimethoprim Average = 93%, Range |  |
| **Dissolution** |  |  |  | 91 - 94% (RSD=1.5%, n=6) |  |
| **Assay** | HPLC | USP 38 NF 33 page 5399 | 93.0 -107.0% | Sulfamethoxazole 99.3% (RSD=0.6%, n=9) | **COMPLIES** |
| **Assay** |  |  |  | Trimethoprim 93.1% (RSD=0.4%, n=9) |  |

**CONCLUSION:** The sample complies with the specifications for the tests performed.

|  |  |  |  |
| --- | --- | --- | --- |
| **ANALYST** | DR. LORNA WANGARI | .......................................................... | DATE: 14-09-2016 |
| **REVIEWER** | DR. GEORGE WANG'ANG'A | .......................................................... | DATE: 26-09-2016 |
| **ANALYST:** | ERNEST MBAE | .......................................................... | DATE: 09-26-2016 |
| **DIRECTOR:** | HEZEKIAH CHEPKWONY | .......................................................... | DATE: 09-27-2016 |