**CERTIFICATE OF ANALYSIS**

**CERTIFICATE No:**  CAN/2016-17/XYZ

|  |  |  |  |
| --- | --- | --- | --- |
| **PRODUCT** | LOKMALÃƒâ€šÃ‚Â® SUPENSION | | REF. NO: NDQD201508176 |
| **DATE RECEIVED:** 24.08.2015 | **LABEL CLAIM:** | Each 60 mL suspension after reconstitution contains: Artemether 180 mg, Lumefantrine 1080 mg & Excipients q.s. | |
| **BATCH NO.:** 3516T | **PRESENTATION:** | Yellow granules for reconstitution into suspension for oral administration, contained in a 60 mL amber coloured glass bottle, packed packed along with 10 mL clear plastic measuring cap in a unit box. | |
| **MGF. DATE:** Jul. 2014 | **MANUFACTURER:** | MZOR Pharmaceutical Industries Ltd. | |
| **EXP. DATE:** Jul. 2016 | **ADDRESS:** | Plot 3C, Block A, Aswani Market Road, Off Oshodi - Apapa Express Way, Isolo, Lagos NIGERIA. | |
| **CLIENT REF NO:** | **CLIENT:** | Pharmacy Board Of Sierra Leone, Central Medical Stores, Free Town, SIERRA LEONE | |
|  | **TEST(S) REQUESTED:** | Microbial Load, Identification, Assay, Acidity/Alkalinity | |

**RESULTS**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **TEST** | **METHOD** | **COMPEDIA** | **SPECIFICATION** | **DETERMINED** | **REMARKS** |
| **Microbial Load** | Plate Count | BP 2016 Vol. V App. XVI B | Total Aerobic Microbial Count (TAMC) <200 CFU/mL Total Combined Yeasts /Moulds Count (TYMC) <20 CFU/mL | < 10 TAMC CFU Per mL < 10 TYMC CFU Per mL | **COMPLIES** |
| **Identification** | HPLC | Adopted In-house Method | RT of the component Peaks in the assay sample preparation correspond to those in the assay standard preparation | Super-imposable peak at RT 5.4 & 5.6 -/+ 10% min. present in both the assay sample and standard preparations | **COMPLIES** |
| **Assay** | HPLC:HPLC | Adopted In-house Method | 90.0 - 110.0% | Artemether 103.5% (RSD = 1.9%; n = 8) | **COMPLIES:COMPLIES** |
| **Assay** | HPLC:HPLC | Adopted In-house Method | 90.0 - 110.0% | Lumefantrine 91.0% (RSD = 1.5%; n = 8) | **COMPLIES:COMPLIES** |
| **Acidity/Alkalinity** | pH | Adopted In-house Method | 3.0 - 5.0 | 4.7 | **COMPLIES** |

**CONCLUSION:** The product complies with the specifications for the tests performed.

|  |  |  |  |
| --- | --- | --- | --- |
| ANALYST: | MR. F. NAULA |  | DATE:21-12-2015 |
| ANALYST: | DR. E. MUTUA |  | DATE:21-12-2015 |
| ANALYST: | DR. G. WANG'ANG'A |  | DATE:21-12-2015 |
| ANALYST: | ERNEST MBAE |  | DATE:12-23-2015 |
| DIRECTOR: | HEZEKIAH CHEPKWONY |  | DATE:01-15-2016 |