**CERTIFICATE OF ANALYSIS**

**CERTIFICATE No:**  CAN/2016-17/XYZ

|  |  |  |  |
| --- | --- | --- | --- |
| **PRODUCT** | DTO TABLETS | | REF. NO: NDQD201510453 |
| **DATE RECEIVED:** 22.10.2015 | **LABEL CLAIM:** | Each film coated tablet contains Ofloxacin USP 200 mg and Ordinazole 500 mg. | |
| **BATCH NO.:** DDB 1501 | **PRESENTATION:** | Orange coloured, caplet shaped, biconvex faced tablets, single scored on one face and plain on the other, packed in a blister pack of 10 tablets and 1 such strip in a unit box. | |
| **MGF. DATE:** May 2015 | **MANUFACTURER:** | CORAL Laboratories Ltd. | |
| **EXP. DATE:**  Apr. 2018 | **ADDRESS:** | Plot No.57/1 (16) Brenslore Dunetha, Nani Doman -396210, INDIA. | |
| **CLIENT REF NO:** | **CLIENT:** | TWOKAY CHEMICALS LTD P.O BOX 46169-00100 NAIROBI,KENYA | |
|  | **TEST(S) REQUESTED:** | Uniformity of Weight, Identification, Dissolution, Assay | |

**RESULTS**

|  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| **TEST** | **METHOD** | **COMPEDIA** | **SPECIFICATION** | **DETERMINED** | **REMARKS** |
| **Uniformity of Weight** | Weight | BP 2016 Vol. V App XII C | NMT 2 tablets deviate by more than 5% from mean weight | None Deviates | **COMPLIES** |
| **Identification** | HPLC | Adopted In - House Method | RT of the component peaks in the assay sample preparation correspond to those in the assay standard preparation | Super-imposable peak at RT 3.3 and 4.9 -/+ 10% min. present in both the assay sample and standard preparations | **COMPLIES** |
| **Dissolution** | HPLC | Adopted In - House Method | No tablets less than 80% (n=6 | Oflocaxin 99.6% (RSD=3.88%, n=6) | **COMPLIES:COMPLIES** |
| **Dissolution** | HPLC | Adopted In - House Method | No tablets less than 80% (n=6 | Ornidazole 98.1% (RSD=4.0%, n=6) | **COMPLIES:COMPLIES** |
| **Assay** | HPLC:HPLC | Adopted In - House Method | 90.0 -110.0% | Oflocaxin 100.2% (RSD=1.2%, n=9) | **COMPLIES:COMPLIES** |
| **Assay** | HPLC:HPLC | Adopted In - House Method | 90.0 -110.0% | Ornidazole 102.6% (RSD=1.7%, n=9) | **COMPLIES:COMPLIES** |

**CONCLUSION:** The product complies with the specifications for the tests performed.

|  |  |  |  |
| --- | --- | --- | --- |
| ANALYST: | DR. L. KARANJA |  | DATE:04-02-2016 |
| ANALYST: | DR. E. TANUI |  | DATE:12-02-2016 |
| ANALYST: | ERNEST MBAE |  | DATE:02-15-2016 |
| DIRECTOR: | HEZEKIAH CHEPKWONY |  | DATE:02-22-2016 |